LITERATURE SYNTHESIS

Accuracy, Calibration, Autonomy, and Physical Integration of Low-Cost Outdoor Sensor Boxes

Summary of the 19 research sources currently included in the Enclosure-Research repository

| Prepared for | Principal Investigator review |
| --- | --- |
| Working paper | Calibration, Ambient Conditions, and Integration Effects on the Accuracy and Autonomy of a Low-Cost Outdoor Multi-Sensor Box |
| Date | June 11, 2026 |

Purpose

This report connects each literature source to the proposed study, summarizes the reported sensors, materials, and geometries, and identifies evidence-supported considerations for the next experimental phase.

# Executive Summary

The literature supports framing the lab's sensor box as an integrated measurement instrument rather than as a collection of individual sensors or an enclosure-only design. Field accuracy depends on calibration, but it also depends on airflow, direct and reflected radiation, sensor placement, electronics self-heating, moisture protection, power, communication, storage, and maintenance.

The strongest recurring result is that passive shields fail most noticeably during high-solar, low-wind conditions. Stacked plates, overlapping louvers, reflective surfaces, and separated electronics improve performance without consuming power. Active aspiration can further reduce airflow-dependent bias, but the fan directly reduces autonomy and adds a failure mode. Long-duration deployments consistently show that connectors, mounts, batteries, charging, storage, communications, and firmware may fail before the sensing element itself.

## Implications for the Proposed Paper

Begin with the current lab box as an unchanged baseline and compare it with reference instruments.

Separate the sealed electronics/power region from sensors that require ambient airflow.

Evaluate raw and calibrated accuracy, but do not use calibration to conceal avoidable enclosure heating or airflow problems.

Report autonomy using runtime before intervention, expected versus received samples, uptime, maintenance events, and failure types.

Use a small, justified design comparison: current box, strong passive shield, and optional aspirated benchmark.

Treat material, color/finish, geometry, airflow, sealing, and sensor-to-electronics distance as measurable factors.

## Cross-Paper Findings

Sensor selection

Digital temperature/RH sensors are compact and easy to integrate, but their readings are strongly affected by shielding, airflow, condensation, and placement.

Optical PM and electrochemical gas sensors require environmental correction, repeated calibration, and controlled exposure to ambient air.

Reference-grade Pt100 studies provide strong shield-comparison methods, even when the final lab sensor is lower cost.

Material and geometry

White/off-white ASA is the best-supported printed outdoor material; PLA is useful for prototypes but requires durability caution.

Reflective foil and white corrugated plastic provide very low-cost geometry benchmarks.

Passive stacked plates generally outperform open or weakly ventilated tube geometries.

IP-rated boxes are appropriate for electronics, but ambient sensors should be external or in a separately ventilated region.

Autonomy and reliability

Battery life alone is not autonomy; physical intervention can be triggered by water, corrosion, memory, fan, firmware, RF, connector, or mounting failures.

Pre/post co-location and device-health logging improve the ability to identify drift and limiting failures.

# Literature at a Glance

The table below identifies the primary contribution of each source. Detailed source profiles follow.

| # | Paper | Primary contribution | Most relevant design issue |
| --- | --- | --- | --- |
| 1 | Theisen et al. (2020) | Complete system / autonomy | The sensing element is only one part of field accuracy and lifetime. Integration and maintenance failures can dominate. |
| 2 | Tatsumi et al. (2021) | Complete system / autonomy | Low-cost, modular hardware is practical, but exact material reporting and external-connection protection are necessary for reproducibility. |
| 3 | Grimsley et al. (2022) | Complete system / autonomy | A modular sealed logger can support long deployments when sensing, power, communications, and maintenance are designed together. |
| 4 | Daepp et al. (2022) | Complete system / autonomy | Completeness and remote serviceability are central measures of whether a low-cost box is useful in practice. |
| 5 | Lazarescu (2015) | Complete system / autonomy | Theoretical battery life does not establish autonomy; observed mechanical, thermal, RF, and firmware behavior does. |
| 6 | Botero-Valencia et al. (2022) | Shield geometry / material | A modular passive cone-stack shield is a defensible low-cost baseline, but material and color were not fully isolated. |
| 7 | Tarara and Hoheisel (2007) | Shield geometry / material | Radiation blocking alone is insufficient: reflected radiation and ventilation area determine passive-shield error. |
| 8 | Holden et al. (2013) | Shield geometry / material | Simple, inexpensive geometry can be effective enough for distributed sensing, but passive limitations remain. |
| 9 | Liu et al. (2023) | Shield geometry / material | Mixed surface finishes and guided natural ventilation offer performance gains, with added fabrication and coating-maintenance complexity. |
| 10 | García Izquierdo et al. (2024) | Shield geometry / material | Long-duration comparison must evaluate operational reliability as well as thermometer error. |
| 11 | Deford et al. (2025) | Shield geometry / material | Active airflow is a useful accuracy benchmark and a direct test of the accuracy-autonomy tradeoff. |
| 12 | Jin et al. (2026) | Shield geometry / material | Advanced passive geometry may reduce error without fan power, but complexity and weathering must be assessed. |
| 13 | deSouza et al. (2022) | Calibration / drift | Calibration performance must be tested beyond the location and period used to fit the model. |
| 14 | Clements et al. (2017) | Calibration / drift | Low-cost sensor boxes need a defined quality-assurance workflow rather than a one-time validation. |
| 15 | Giordano et al. (2021) | Calibration / drift | PM calibration and enclosure design are deployment-specific and cannot be treated independently. |
| 16 | Vajs et al. (2021) | Calibration / drift | Interpretable linear and multiple-linear models provide a defensible baseline before more complex methods. |
| 17 | Schneider et al. (2023) | Calibration / drift | Pre- and post-deployment co-location improves detection of drift and field-induced change. |
| 18 | Diez et al. (2024) | Calibration / drift | Long-term validation requires documentation of hardware, firmware, processing, location, and uncertainty. |
| 19 | Winter et al. (2025) | Calibration / drift | Recalibration interval and physical-maintenance interval are separate engineering quantities. |

# Complete Sensor-Box and Field-System Studies

These papers show how sensors, packaging, power, communication, and maintenance behave as one deployed system.

## 1. Theisen et al. (2020)

More Science with Less: Evaluation of a 3D-Printed Weather Station

DOI: https://doi.org/10.5194/amt-13-4699-2020

| Sensors | MCP9808; HTU21D; BMP280; SI1145; Hall-effect wind, direction, and rain sensors. |
| --- | --- |
| Materials | Off-white/gray ASA; PVC frame; electrical junction box; polyurethane, PTFE filter, conformal coating, and silicone. |
| Geometry | Integrated 3D-PAWS station; naturally aspirated shield; crossbar-mounted wind/UV sensors; separate junction box for logger. |
| Experimental setup | Eight-month outdoor deployment beside an Oklahoma Mesonet reference station. |

What it shows: Several low-cost measurements can be useful, but wind dependence, corrosion, moisture, connectors, printed fasteners, and moving parts strongly affect total system performance.

Brief conclusion: The sensing element is only one part of field accuracy and lifetime. Integration and maintenance failures can dominate.

Importance to this research: Provides the closest precedent for evaluating the lab box as a complete deployed instrument and for recording both accuracy and failure events.

Things to consider moving forward

Log internal box temperature and wind speed so radiation-shield performance can be separated from sensor calibration.

Use metal structural fasteners and explicitly inspect connectors, drain paths, coatings, and exposed printed parts.

## 2. Tatsumi et al. (2021)

An Open-Source, Low-Cost Measurement System for Collecting Hydrometeorological Data in the Open Field

DOI: https://doi.org/10.3390/technologies9040078

| Sensors | LM60BIZ air/soil temperature; YL-69 soil moisture; G2711-01 GaAsP photodiode; GPS and micro-SD. |
| --- | --- |
| Materials | 3D-printed protective case using unspecified resin; handmade photodiode cover. |
| Geometry | Compact rectangular logger case with external sensors connected through detachable terminals; solar-powered system. |
| Experimental setup | Grass-field deployment with reference comparisons; 30-day temperature and 19-day soil-moisture/PPFD measurements. |

What it shows: An open-field system must be evaluated as a package of sensing, power, storage, enclosure, and external connections.

Brief conclusion: Low-cost, modular hardware is practical, but exact material reporting and external-connection protection are necessary for reproducibility.

Importance to this research: Supports the baseline inventory and reinforces documenting the complete lab box before modification.

Things to consider moving forward

Record exact material chemistry rather than only stating that a case is 3D printed.

Include connector sealing, strain relief, and external-probe replacement in maintainability testing.

## 3. Grimsley et al. (2022)

Experiences in LP-IoT: EnviSense Deployment of Remotely Reprogrammable Environmental Sensors

DOI: https://doi.org/10.1145/3477085.3478988

| Sensors | MaxBotix ultrasonic range sensor or SDI-12 pressure transducer; internal T/RH/pressure; battery voltage/current; GPS; LoRa. |
| --- | --- |
| Materials | Commercial IP67 enclosure; exact enclosure material not reported. |
| Geometry | Sealed rectangular logger with four sealed ports and external measurement probes. |
| Experimental setup | More than 12 nodes at two sites; six months of continuous field measurement discussed. |

What it shows: Sealed electronics, external probes, remote management, and device-health sensing improve autonomy and diagnosability.

Brief conclusion: A modular sealed logger can support long deployments when sensing, power, communications, and maintenance are designed together.

Importance to this research: Strong precedent for a two-zone architecture and for treating battery and communication health as measurement-system outputs.

Things to consider moving forward

Separate protected electronics from ambient-air sensors.

Log battery, resets, and communication health alongside environmental data.

## 4. Daepp et al. (2022)

Eclipse: An End-to-End Platform for Low-Cost, Hyperlocal Environmental Sensing in Cities

DOI: https://doi.org/10.1109/IPSN54338.2022.00010

| Sensors | PM2.5; temperature; RH; pressure; optional O3, NO2, SO2, and CO modules. |
| --- | --- |
| Materials | Custom weatherproof enclosure; exact material and airflow path not reported. |
| Geometry | Modular solar-powered, cellular-connected urban sensor box. |
| Experimental setup | 115 devices deployed across Chicago; more than 90% of expected sensor-hours collected during the reported period. |

What it shows: Large-scale deployment is possible when the platform is modular and expected sensor-hours are treated as an operational metric.

Brief conclusion: Completeness and remote serviceability are central measures of whether a low-cost box is useful in practice.

Importance to this research: Provides a clear autonomy metric and supports designing reset/service actions that do not require opening the box.

Things to consider moving forward

Report expected versus received sensor-hours.

Evaluate site-specific mounting and solar-panel heating effects.

## 5. Lazarescu (2015)

Design and Field Test of a WSN Platform Prototype for Long-Term Environmental Monitoring

DOI: https://doi.org/10.3390/s150409481

| Sensors | NTC temperature transducer; low-power radio; battery and system-health measurements. |
| --- | --- |
| Materials | Protective resin; compact node package; sealed plastic gateway package inside a divided wooden housing. |
| Geometry | Downward-facing vented node; solar-powered gateway with separated compartments. |
| Experimental setup | Fifty sensor nodes and two gateways deployed for approximately 2.5 months in an olive field. |

What it shows: Solar heating, solder fatigue, antenna placement, charger failure, RF attenuation, and faulty-node traffic can limit autonomy.

Brief conclusion: Theoretical battery life does not establish autonomy; observed mechanical, thermal, RF, and firmware behavior does.

Importance to this research: Directly informs the paper's failure-mode and maintenance-event framework.

Things to consider moving forward

Test antenna clearance and enclosure effects on communication.

Include heartbeats, self-test, battery-state validation, and failure propagation in the deployment plan.

# Radiation-Shield Geometry and Material Studies

These papers provide the strongest evidence for selecting and comparing passive and active shield designs.

## 6. Botero-Valencia et al. (2022)

Design and Implementation of 3-D Printed Radiation Shields for Environmental Sensors

DOI: https://doi.org/10.1016/j.ohx.2022.e00267

| Sensors | Three identical SHT10 temperature/RH sensors; Vantage Pro reference; Particle Argon logger. |
| --- | --- |
| Materials | Orange PLA and white ASA; PVC mount; metal spacers; IP69 electronics box. |
| Geometry | Six truncated cones: five perforated cones and a solid roof, separated by 20 mm spacers. |
| Experimental setup | Exposed, PLA-shielded, and ASA-shielded sensors compared outdoors, plus 30/90-day weathering tests. |

What it shows: Both printed shields reduced radiation error; ASA was favored for outdoor UV resistance and durability.

Brief conclusion: A modular passive cone-stack shield is a defensible low-cost baseline, but material and color were not fully isolated.

Importance to this research: Provides the most direct literature basis for testing printed passive shields and comparing material/weathering behavior.

Things to consider moving forward

Control color and surface finish when comparing base materials.

Treat cone count, spacing, and open area as independent geometry variables.

## 7. Tarara and Hoheisel (2007)

Low-Cost Shielding to Minimize Radiation Errors of Temperature Sensors in the Field

DOI: https://doi.org/10.21273/HORTSCI.42.6.1372

| Sensors | Same temperature-sensor type used across test and comparison shields. |
| --- | --- |
| Materials | Reflective foil insulation with high albedo and low emissivity. |
| Geometry | Plate, stacked-plate, cone, double-cone, tube, rocket, pagoda, and perforated-tube variants. |
| Experimental setup | Controlled outdoor comparison under changing solar irradiance and wind. |

What it shows: Stacked plates performed best; open bottoms admitted reflected radiation; tube shields needed substantial sidewall perforation.

Brief conclusion: Radiation blocking alone is insufficient: reflected radiation and ventilation area determine passive-shield error.

Importance to this research: Provides the clearest evidence for selecting and parameterizing passive geometry variants.

Things to consider moving forward

Measure vent/open-area ratio and include reflected ground radiation.

Avoid open-bottom and weakly perforated tube designs.

## 8. Holden et al. (2013)

Design and Evaluation of an Inexpensive Radiation Shield for Monitoring Surface Air Temperatures

DOI: https://doi.org/10.1016/j.agrformet.2013.06.011

| Sensors | Low-cost temperature sensors, including LogTag units. |
| --- | --- |
| Materials | White corrugated-plastic sheet. |
| Geometry | Folded Gill-style series of plates. |
| Experimental setup | Three-site outdoor comparison against commercial passive shields, RAWS, and aspirated ASOS reference. |

What it shows: A roughly USD 3 shield performed similarly to a commercial passive shield, while low-wind full-sun bias remained.

Brief conclusion: Simple, inexpensive geometry can be effective enough for distributed sensing, but passive limitations remain.

Importance to this research: Creates a useful low-cost manufacturability benchmark for the lab's enclosure comparisons.

Things to consider moving forward

Compare build repeatability, weathering, and cost alongside accuracy.

Use an aspirated or higher-grade reference to quantify remaining passive bias.

## 9. Liu et al. (2023)

Design and Experiments of a Naturally-Ventilated Radiation Shield for Ground Temperature Measurement

DOI: https://doi.org/10.3390/atmos14030523

| Sensors | Pt100 platinum resistance thermometer. |
| --- | --- |
| Materials | Silver-coated aluminum outer plates; black inner surfaces; white-resin internal parts. |
| Geometry | Multilayer shield with two square plates, annular layers, diversion pipe, eight vents, and truncated-cone reflector. |
| Experimental setup | CFD-informed design followed by outdoor experimental validation. |

What it shows: Outer reflectivity, inner surface treatment, venting, and lower reflectors can address direct and reflected radiation together.

Brief conclusion: Mixed surface finishes and guided natural ventilation offer performance gains, with added fabrication and coating-maintenance complexity.

Importance to this research: Expands the paper beyond polymer choice to include surface finish and reflected-radiation management.

Things to consider moving forward

Treat exterior reflectance and interior finish as separate variables.

Include vent count/area and coating degradation in the decision framework.

## 10. García Izquierdo et al. (2024)

COAT Project: Intercomparison of Thermometer Radiation Shields in the Arctic

DOI: https://doi.org/10.3390/atmos15070841

| Sensors | Forty-one calibrated four-wire Pt100 thermometers; ultrasonic anemometers and reference meteorology. |
| --- | --- |
| Materials | Ten commercial shield models; detailed proprietary materials vary. |
| Geometry | Natural and active shields, including multiplate and Stevenson-screen configurations. |
| Experimental setup | Fourteen-month Arctic intercomparison with pre/post calibration. |

What it shows: Shield differences depend strongly on wind and radiation; fan, logger, power, memory, mounting, and storm failures caused data loss.

Brief conclusion: Long-duration comparison must evaluate operational reliability as well as thermometer error.

Importance to this research: Provides a strong precedent for long-term failure logging and passive-versus-active comparison.

Things to consider moving forward

Log fan operation, storage capacity, power status, and mount condition.

Use redundant probes where possible to distinguish shield effects from sensor failure.

## 11. Deford et al. (2025)

Development of a Low-Cost, 3D-Printed, Aspirated Air Temperature Measurement Radiation Shield

DOI: https://doi.org/10.1175/JTECH-D-24-0006.1

| Sensors | Compact air-temperature sensor; Sunon 5 V fan. |
| --- | --- |
| Materials | White PLA; hot glue during testing; fan and exposed circuits. |
| Geometry | Three-part double-wall 90-degree elbow with insect screen, central supports, and angled fan diffuser. |
| Experimental setup | Outdoor comparison with passive and commercial aspirated reference shields. |

What it shows: Low-cost aspiration can achieve strong accuracy, but the fan consumes 680 mW and creates durability and maintenance risks.

Brief conclusion: Active airflow is a useful accuracy benchmark and a direct test of the accuracy-autonomy tradeoff.

Importance to this research: Provides the clearest candidate active design for comparison with passive lab-box options.

Things to consider moving forward

Measure fan power, flow, and failure behavior.

Replace temporary PLA/glue/exposed circuitry with weatherable materials and permanent assembly for long deployment.

## 12. Jin et al. (2026)

Design and Experimental Validation of a High-Accuracy Naturally Ventilated Radiation Shield

DOI: https://doi.org/10.3390/atmos17030272

| Sensors | Thin-film Pt100 in an 8 mm copper spherical shell. |
| --- | --- |
| Materials | CFD comparison of plastic, wood, Fe-Ni alloy, and aluminum; mirror-finish aluminum considered. |
| Geometry | Two square shading plates with symmetric bowl-cover flow-guiding shrouds and 30 mm spacing. |
| Experimental setup | CFD material/geometry optimization followed by outdoor validation. |

What it shows: Material thermal properties and guided passive airflow jointly affect radiation error.

Brief conclusion: Advanced passive geometry may reduce error without fan power, but complexity and weathering must be assessed.

Importance to this research: Supports evaluating airflow guidance and upward/reflected radiation rather than only plate count.

Things to consider moving forward

Balance modeled accuracy against manufacturability, size, mounting, and coating durability.

Validate CFD-derived designs across multiple weather states.

# Calibration, Drift, and Air-Quality Sensor Studies

These papers define calibration, drift, environmental correction, and long-term validation requirements.

## 13. deSouza et al. (2022)

Calibrating Networks of Low-Cost Air Quality Sensors

DOI: https://doi.org/10.5194/amt-15-6309-2022

| Sensors | Canary-S boxes using Plantower PMS5003 PM sensing with integrated temperature/RH. |
| --- | --- |
| Materials | Commercial outdoor box; material and internal airflow not reported. |
| Geometry | Commercial networked box; detailed inlet/outlet and internal layout not reported. |
| Experimental setup | Twenty-four school deployments; seven boxes co-located with five FEM reference sites. |

What it shows: Calibration evaluated at one co-location may fail to transfer across time or space.

Brief conclusion: Calibration performance must be tested beyond the location and period used to fit the model.

Importance to this research: Directly supports time-separated and, if possible, site-separated validation of the lab box.

Things to consider moving forward

Document PM inlet, fan path, sensor placement, and enclosure material.

Report calibration transfer performance rather than only training/co-location error.

## 14. Clements et al. (2017)

Low-Cost Air Quality Monitoring Tools: From Research to Practice

DOI: https://doi.org/10.3390/s17112478

| Sensors | Optical PM, electrochemical gas, metal-oxide gas, and multi-sensor systems discussed generally. |
| --- | --- |
| Materials | No single enclosure material reported. |
| Geometry | No single geometry; air exchange and internal heat are identified as important. |
| Experimental setup | Workshop summary synthesizing laboratory, co-location, mobile, community, and network experience. |

What it shows: Co-location, repeated calibration, drift checks, and environmental measurements are common requirements.

Brief conclusion: Low-cost sensor boxes need a defined quality-assurance workflow rather than a one-time validation.

Importance to this research: Provides broad best-practice justification for calibration and enclosure-condition logging.

Things to consider moving forward

Measure internal heat and air exchange rather than treating the box as neutral.

Place environmental-correction sensors close to pollutant sensors while exposing them to ambient air.

## 15. Giordano et al. (2021)

From Low-Cost Sensors to High-Quality Data

DOI: https://doi.org/10.1016/j.jaerosci.2021.105833

| Sensors | Low-cost optical particulate-matter sensors as a class. |
| --- | --- |
| Materials | No single enclosure material reported. |
| Geometry | No single enclosure; packaging and airflow identified as influences. |
| Experimental setup | Review of laboratory evaluation, co-location, and network deployment studies. |

What it shows: PM response depends on humidity, aerosol properties, sensor age, site, and packaging/airflow.

Brief conclusion: PM calibration and enclosure design are deployment-specific and cannot be treated independently.

Importance to this research: Defines environmental covariates and limitations that must be considered if the lab box includes PM sensing.

Things to consider moving forward

Include humidity, inlet airflow, aging, and local aerosol conditions in PM analysis.

Do not infer a universally transferable calibration.

## 16. Vajs et al. (2021)

Developing Relative Humidity and Temperature Corrections for Low-Cost Sensors Using Machine Learning

DOI: https://doi.org/10.3390/s21103338

| Sensors | Alphasense gas cells; Bosch BME280; Plantower PMS7003. |
| --- | --- |
| Materials | Commercial AQ10x enclosure; material and airflow not reported. |
| Geometry | Commercial outdoor air-quality station; internal layout not reported. |
| Experimental setup | One station co-located with a Serbian public monitoring station; LR, MLR, and ML corrections compared. |

What it shows: Temperature and RH correction can improve low-cost gas/PM measurements, but sensor-specific sensitivity remains.

Brief conclusion: Interpretable linear and multiple-linear models provide a defensible baseline before more complex methods.

Importance to this research: Directly supports the proposed calibration sequence for the lab box.

Things to consider moving forward

Start with linear and multiple-linear calibration and use a time-separated test period.

Ensure T/RH sensors measure the air reaching the pollutant sensors, not only the electronics-box microclimate.

## 17. Schneider et al. (2023)

Deployment and Evaluation of a Network of Open Low-Cost Air Quality Sensor Systems

DOI: https://doi.org/10.3390/atmos14030540

| Sensors | Alphasense NO2, CO, NO, O3; ELT D-300 CO2; OPC-N3 PM; T/RH/pressure. |
| --- | --- |
| Materials | AirSensEUR boxes; enclosure material not reported. |
| Geometry | Complete boxes co-located together, then distributed to roadside sites. |
| Experimental setup | Five-week pre-deployment co-location, roadside deployment, and post-deployment evaluation. |

What it shows: High RH affected optical PM; unit-to-unit differences, drift, and failures required pre/post deployment evaluation.

Brief conclusion: Pre- and post-deployment co-location improves detection of drift and field-induced change.

Importance to this research: Supports comparing multiple lab boxes and checking calibration after deployment.

Things to consider moving forward

Perform pre/post co-location if deployment duration permits.

Treat humidity, unit variability, and multi-sensor maintenance as primary factors.

## 18. Diez et al. (2024)

Long-Term Evaluation of Commercial Air Quality Sensors: The QUANT Study

DOI: https://doi.org/10.5194/amt-17-3809-2024

| Sensors | Forty-three commercial devices containing 119 gas and 118 PM sensors. |
| --- | --- |
| Materials | Commercial protective cases; materials and internal layouts generally proprietary. |
| Geometry | Multiple commercial black-box systems. |
| Experimental setup | Three-year evaluation across UK urban environments, including relocation and uncertainty analysis. |

What it shows: Performance varied substantially between systems; relocation and manufacturer correction changes affected results.

Brief conclusion: Long-term validation requires documentation of hardware, firmware, processing, location, and uncertainty.

Importance to this research: Supports treating the lab box as a transparent system and evaluating more than RMSE alone.

Things to consider moving forward

Document every hardware, firmware, and calibration change.

Evaluate inter-device precision, uncertainty, relocation, and long-term stability.

## 19. Winter et al. (2025)

Sustained Performance of Low-Cost Air Quality Sensors in Long-Term Deployments

DOI: https://doi.org/10.1021/acssensors.5c00566

| Sensors | Alphasense NO2-B43F, Ox-B431, NO-B4, and CO-B4 electrochemical sensors. |
| --- | --- |
| Materials | BEACO2N node enclosure not described in detail. |
| Geometry | Dense-network node packaging; geometry not the focus. |
| Experimental setup | Three-year San Francisco Bay Area network evaluation. |

What it shows: Calibrated electrochemical performance can remain stable for years, while mechanical/electrical failures may occur first.

Brief conclusion: Recalibration interval and physical-maintenance interval are separate engineering quantities.

Importance to this research: Strengthens the paper's focus on identifying the true limiting factor rather than assuming sensor drift dominates.

Things to consider moving forward

Track recalibration need separately from physical interventions.

Record board, connector, power, and enclosure failures alongside sensing-element behavior.

# Recommended Experimental Framing

The literature supports an engineering comparison centered on the lab's existing box. The primary result should be identification of the limiting factor: sensor calibration, physical integration, power, firmware/data handling, or maintenance.

## Recommended Initial Designs

| Design | Purpose |
| --- | --- |
| Current lab box | Unmodified baseline that establishes present accuracy, autonomy, and failure modes. |
| Strong passive option | White reflective stacked-plate or truncated-cone shield with electronics isolated in an IP-rated compartment. |
| Active benchmark | Aspirated shield used only if the power budget permits; records fan power and fan-operating status. |

## Measurements Needed

Reference-matched sensor readings and timestamps.

Solar exposure or radiation proxy, wind, rain, ambient temperature, and RH.

Battery voltage/current, fan state if used, resets, storage and communication health.

Expected and received sample counts by sensor channel.

Maintenance, calibration, physical inspection, and failure-event log.

Material, color/finish, geometry dimensions, vent/open-area ratio, and sensor placement.

## Questions to Resolve Before Deployment

What variables must the lab box measure, and what reference instruments are available for each?

What calibrated-error, completeness, and autonomy thresholds define a useful deployment?

Which current component is most likely to limit performance: sensor, enclosure, power, firmware, or maintenance?

Can the lab fabricate ASA reliably, or should PETG/coated PETG and purchased IP-rated boxes be the practical candidates?

Is active aspiration justified by the expected accuracy improvement and available power budget?

How long must the deployment run to capture representative solar, wind, humidity, rain, and temperature conditions?

# Overall Conclusion

Across the 19 sources, the strongest evidence supports a two-zone outdoor sensor-box architecture: protected electronics and power in a sealed, serviceable enclosure, with ambient sensors placed in an external or separately ventilated shield. A light-colored passive stacked-plate or cone shield is the most defensible low-power baseline. Active aspiration is valuable as an accuracy benchmark but must be evaluated against its power and maintenance cost. The paper should therefore report calibration improvement and physical-design performance together, while using observed intervention-free runtime and data completeness to define autonomy.

The final contribution should be a practical decision framework showing which design factors most affect the lab's sensor box and which changes provide measurable improvement without unnecessary complexity.

# Reference List

Theisen et al. (2020). More Science with Less: Evaluation of a 3D-Printed Weather Station. https://doi.org/10.5194/amt-13-4699-2020

Tatsumi et al. (2021). An Open-Source, Low-Cost Measurement System for Collecting Hydrometeorological Data in the Open Field. https://doi.org/10.3390/technologies9040078

Grimsley et al. (2022). Experiences in LP-IoT: EnviSense Deployment of Remotely Reprogrammable Environmental Sensors. https://doi.org/10.1145/3477085.3478988

Daepp et al. (2022). Eclipse: An End-to-End Platform for Low-Cost, Hyperlocal Environmental Sensing in Cities. https://doi.org/10.1109/IPSN54338.2022.00010

Lazarescu (2015). Design and Field Test of a WSN Platform Prototype for Long-Term Environmental Monitoring. https://doi.org/10.3390/s150409481

Botero-Valencia et al. (2022). Design and Implementation of 3-D Printed Radiation Shields for Environmental Sensors. https://doi.org/10.1016/j.ohx.2022.e00267

Tarara and Hoheisel (2007). Low-Cost Shielding to Minimize Radiation Errors of Temperature Sensors in the Field. https://doi.org/10.21273/HORTSCI.42.6.1372

Holden et al. (2013). Design and Evaluation of an Inexpensive Radiation Shield for Monitoring Surface Air Temperatures. https://doi.org/10.1016/j.agrformet.2013.06.011

Liu et al. (2023). Design and Experiments of a Naturally-Ventilated Radiation Shield for Ground Temperature Measurement. https://doi.org/10.3390/atmos14030523

García Izquierdo et al. (2024). COAT Project: Intercomparison of Thermometer Radiation Shields in the Arctic. https://doi.org/10.3390/atmos15070841

Deford et al. (2025). Development of a Low-Cost, 3D-Printed, Aspirated Air Temperature Measurement Radiation Shield. https://doi.org/10.1175/JTECH-D-24-0006.1

Jin et al. (2026). Design and Experimental Validation of a High-Accuracy Naturally Ventilated Radiation Shield. https://doi.org/10.3390/atmos17030272

deSouza et al. (2022). Calibrating Networks of Low-Cost Air Quality Sensors. https://doi.org/10.5194/amt-15-6309-2022

Clements et al. (2017). Low-Cost Air Quality Monitoring Tools: From Research to Practice. https://doi.org/10.3390/s17112478

Giordano et al. (2021). From Low-Cost Sensors to High-Quality Data. https://doi.org/10.1016/j.jaerosci.2021.105833

Vajs et al. (2021). Developing Relative Humidity and Temperature Corrections for Low-Cost Sensors Using Machine Learning. https://doi.org/10.3390/s21103338

Schneider et al. (2023). Deployment and Evaluation of a Network of Open Low-Cost Air Quality Sensor Systems. https://doi.org/10.3390/atmos14030540

Diez et al. (2024). Long-Term Evaluation of Commercial Air Quality Sensors: The QUANT Study. https://doi.org/10.5194/amt-17-3809-2024

Winter et al. (2025). Sustained Performance of Low-Cost Air Quality Sensors in Long-Term Deployments. https://doi.org/10.1021/acssensors.5c00566